## Slide 1: Chem. 133 – 2/16 Lecture

## Slide 2: Announcements
HW1.2 problems due
Quiz 2 today
Today’s Lecture
  Noise

## Slide 3: Signal Averaging
If the noise is random and well known, repeat measurements improve S/N because signal is additive while noise adds as (n)0.5 (based on propagation of uncertainty rules)
Note: in some cases, averaging can affect qualitative information as well as quantitative information (e.g. mass spectrometer measured mass)
(S/N)n = [(S/N)n=1]n/(n)0.5 = [(S/N)n=1](n)0.5

## Slide 4: Signal Averaging - Question
A 1H NMR is performed on a small amount of sample expected to be the compound at right:
With 16 scans the S/N observed for the c 1H peak is 17.
How many scans are needed so that the minimum peak has a S/N of 3? (Assume all peaks have the same width)
a
b
c

## Slide 5: Signal Averaging Another Example
In mass spectrometry, and in particular with time-of-flight mass spectrometers, mass measurement is measured on many ions
Instrument resolution is good, but insufficient for high resolution on single measurements (resolution of 15,000 gives s ~ 0.1 amu for 1344 peak)
To meet “accurate mass” requirement, errors less than 5 ppm (0.007 amu) are required.
2s ~ 0.2 amu
x axis is mass
A single measurement will never meet high resolution requirement, but averaging will result in an improved average value (mean mass).
For n > 50, 95% CI becomes mean + 1.96s/(n)0.5 or to reach 0.007 amu, would require roughly 784 “counts” or individual measurements

## Slide 6: Noise Sources – Fundamental Types
Thermal Noise = Johnson Noise (voltage associated)
- where:
kB = Boltzmann’s constant, T = temp. (K), R = resistance (W), and B = bandwidth (Hz) = range of frequencies accepted
- Solutions: cool devices, use lower R values, reduce bandwidth
B. Shot noise (current associated)
- Solutions: reduce bandwidth, use internally amplified transducers
where q= fundamental charge = 1.6 x 10-19 C and I = current

## Slide 7: Noise Sources – Other Types
Flicker Noise (or 1/f noise or pink noise)
  Occurs at low frequencies
  Can result from environmental changes (e.g. change in light intensity over time, change in temperature)
  Can be reduced through modulating source

## Slide 8: Noise Flicker Noise Example
lamp
chopper (alternatively reflects light or lets light through)
light detector
sample cell
blank cell
mirrors
Example of equipment for noise reduction
To Digitizer
high pass filter
rectifier
Notes: 8

## Slide 9: Noise Flicker Noise Example: Signals
light detector signal
slow increase in noise over 1st ~100 s
RC Filter only
low f noise removed
RC Filter + diode
Signal following digital filtration
Notes: 9

## Slide 10: Noise Sources – Other Types
B. Interference
  Noise originating from other electrical signals (especially ones that use more power)
  Examples: 60 Hz from power lines, spikes from solenoids, turning heaters on/off
  Solution to problems: 1) use shielded cables, 2) shield major power sources, 3) use differential amplifiers

## Slide 11: Noise Processing to Reduce
Both analog and digital means can be used to reduce noise
Band width reduction can reduce noise
Low pass filters (e.g. RC filters) reduce high frequency noise but lose high frequency signals
Similar methods (e.g. moving averages) for removing high frequency noise can easily be done on digital data
A separate way to reduce noise is to signal average (e.g. collection and averaging spectra), provided experiment can be replicated closely

## Slide 12: Noise Processing to Reduce
Example of moving average (0.5, 2, 10, and 30 s) to remove high frequency noise
Raw Data (collected at 10 Hz)
Raw + 0.5 s data
5 0.1 s points averaged to make 0.5 s point
Only slight reduction of noise observed
Raw + 2 s data
Definite reduction of noise observed
Raw + 10 s data
Even more noise reduction observed, but peak starting to broaden and flatten
Raw + 30 s data
Clearly over filtered (resolution lost by broadening)

## Slide 13: Noise HPLC Example
Type | Noise | Baseline | Signal | Baseline corrected signal | S/N | Peak Width at half height
Unfiltered | 0.002891 | 0.999447 | 1.0519 | 0.052470125 | 18 | 0.202
0.5 s MA | 0.002166 | 0.999461 | 1.0489 | 0.049437891 | 23 | 0.202
2 s MA | 0.001232 | 0.99946 | 1.0458 | 0.046335795 | 38 | 0.2282
10 s MA | 0.000535 | 0.999457 | 1.0414 | 0.041971743 | 79 | 0.255
30s MA | 0.00047 | 0.999512 | 1.0221 | 0.022616734 | 48 | 0.507
Best S/N with 10 s filter
Minimal increase in peak width until 2 s

## Slide 14: Noise Questions
What type of noise is likely to be present when using thermocouples to measure temperature?
Why is modulation normally required to reduce 1/f noise?
What is the percent noise on a current producing transducer which generates signal over a 1000 Hz band if the signal is 10 nA? if the signal is 2.0 pA?
What specific type of noise is reduced best by shielding electronics?
How would use of a low pass filter reduce shot noise?
Suggest one method for reducing thermal noise.
What type of noise is not effectively reduced by using a low pass filter?

## Slide 15: Electronics Additional Questions
Answer the questions 1-3 from the following plots which were obtained from background measurements (instrument noise):
Which plot is most likely shows 1/f noise: ______________________
Which plot when Fourier transformed will produce a plot with a peak at 55 Hz: ______________
If plot c) shows noise from a GC signal in which peaks typically are on the order of 2 s (2000 ms) wide, what can be done to reduce the noise?
a)
b)
c)