## Slide 1: Chem. 133 – 2/21 Lecture

## Slide 2: Announcements
Return Q2 and HW1.2
  I thought Q2 was reasonable – scores were o.k., but not great
  Some parts of homework (e.g. I wanted you to use Excel to solve 1.2.3 and to calculate Sx – the standard uncertainty in the conc. – for each sample) were not clear so will make 2 points bonus (out of 6 points)
Missing Office Hours on Friday
Today’s Lecture
  Noise
  Electrochemistry

## Slide 3: Noise Sources – Other Types
Flicker Noise (or 1/f noise or pink noise)
  Occurs at low frequencies
  Can result from environmental changes (e.g. change in light intensity over time, change in temperature)
  Can be reduced through “modulation” – see next slide

## Slide 4: Noise Flicker Noise Example
lamp
chopper (alternatively reflects light or lets light through)
light detector
sample cell
blank cell
mirrors
Example of equipment (including modulation) for noise reduction
To Digitizer
high pass filter
rectifier
Notes: 4

## Slide 5: Noise Flicker Noise Example: Signals
light detector signal
slow increase in noise over 1st ~100 s
RC Filter only
low f noise removed
RC Filter + diode
Signal following digital filtration
Notes: 5

## Slide 6: Noise Sources – Other Types
B. Interference
  Noise originating from other electrical signals (especially ones that use more power)
  Examples: 60 Hz from power lines, spikes from solenoids, turning heaters on/off
  Solution to problems: 1) use shielded cables, 2) shield major power sources, 3) use differential amplifiers

## Slide 7: Noise Processing to Reduce
Both analog and digital means can be used to reduce noise
Band width reduction can reduce noise
Low pass filters (e.g. RC filters) reduce high frequency noise but lose high frequency signals
Similar methods (e.g. moving averages) for removing high frequency noise can easily be done on digital data
A separate way to reduce noise is to signal average (e.g. collection and averaging spectra), provided experiment can be replicated closely

## Slide 8: Noise Processing to Reduce
Example of moving average (0.5, 2, 10, and 30 s) to remove high frequency noise
Raw Data (collected at 10 Hz)
Raw + 0.5 s data
5 0.1 s points averaged to make 0.5 s point
Only slight reduction of noise observed
Raw + 2 s data
Definite reduction of noise observed
Raw + 10 s data
Even more noise reduction observed, but peak starting to broaden and flatten
Raw + 30 s data
Clearly over filtered (resolution lost by broadening)

## Slide 9: Noise HPLC Example
Type | Noise | Baseline | Signal | Baseline corrected signal | S/N | Peak Width at half height
Unfiltered | 0.002891 | 0.999447 | 1.0519 | 0.052470125 | 18 | 0.202
0.5 s MA | 0.002166 | 0.999461 | 1.0489 | 0.049437891 | 23 | 0.202
2 s MA | 0.001232 | 0.99946 | 1.0458 | 0.046335795 | 38 | 0.2282
10 s MA | 0.000535 | 0.999457 | 1.0414 | 0.041971743 | 79 | 0.255
30s MA | 0.00047 | 0.999512 | 1.0221 | 0.022616734 | 48 | 0.507
Best S/N with 10 s filter
Minimal increase in peak width until 2 s

## Slide 10: Noise Questions
What type of noise is likely to be present when using thermocouples to measure temperature?
Why is modulation normally required to reduce 1/f noise?
What is the percent noise on a current producing transducer which generates signal over a 1000 Hz band if the signal is 10 nA? if the signal is 2.0 pA?
What specific type of noise is reduced best by shielding electronics?
How would use of a low pass filter reduce shot noise?
Suggest one method for reducing thermal noise.
What type of noise is not effectively reduced by using a low pass filter?

## Slide 11: Electronics Additional Questions
Answer the questions 1-3 from the following plots which were obtained from background measurements (instrument noise):
Which plot is most likely shows 1/f noise: ______________________
Which plot when Fourier transformed will produce a plot with a peak at 55 Hz: ______________
If plot c) shows noise from a GC signal in which peaks typically are on the order of 2 s (2000 ms) wide, what can be done to reduce the noise?
a)
b)
c)

## Slide 12: Electrochemistry Overview
Applications
  quantitative analysis
    potential measurement methods (e.g. pH electrode)
    current based measurements (amperometry)
  qualitative analysis (voltammetry)
  note: potential normally gives qualitative information and current quantitative measurements (except with ion selective electrodes/pH meters)
  HPLC/IC detectors
Why Use?
  lower cost (both for instruments and supplies)
  high sensitivity possible (particularly mass sensitivity)
  simpler equipment, more useful for field, in-situ type measurements

## Slide 13: Electrochemistry Redox Reactions
Reduction = loss of charge
  e.g. Fe3+ + e- → Fe2+
Oxidation = gain in charge
  e.g. Pb2+ + 2H2O → PbO2(s) + 4H+ + 2e- (Pb goes from +2 to +4)
Balancing reactions
  review steps in general chemistry book
  example: Zn(s) + Cr2O72- → Zn2+ + Cr3+
  note: based on methods used for problems in this book, full cell balancing may not be needed

## Slide 14: Electrochemistry Fundamental Equations
Relationship between charge, energy and current
  redox reactions involve the exchange of electrons
  when the exchange occurs on an electrode surface, current can be measured
  Total charge transfer = q = nF, where
    n = moles of electrons in reaction and
    F = Faraday’s constant = 96500 C/moles e-
    F = NAvogadro·e (e = elementary charge = 1.6 x 10-19 C)
  Current Produced = I = dq/dt
    or q = ∫I·dt (or = I·t under constant current conditions)
    can be used to determine battery lifetime

## Slide 15: Electrochemistry Fundamental Equations
Relationship between charge, energy and current (continued)
  Electrical work (units = J) = E·q (E = potential in volts and q in C)
  and ΔG = -E·q = -nFE
  under standard conditions (1 M reactant/product conc., 298K, etc.), ΔGº = -nFEº
  ΔGº are given in Tables and allows calculation of K values
  Eº, standard reduction potential, also given in Tables (see Appendix H), but for “half-reactions”

## Slide 16: Electrochemistry Fundamental Equations
Example problem:
  A NiCad battery contains 12.0 g of Cd that is oxidized to Cd(OH)2. How long should the battery last if a motor is drawing 421 mA? Assume 100% efficiency.

## Slide 17: Electrochemistry Galvanic Cells
What are galvanic cells?
  Cells that use chemical reactions to generate electrical energy
  Batteries are examples of useful galvanic cells
  Example reaction
  If reactants are placed in a beaker, only products + heat are produced
  When half reactions are isolated on electrodes, electrical work can be produced
Zn(s) + 2Ag+ → Zn2+ + 2Ag(s)
Salt Bridge
voltmeter
Zn(s)
ZnSO4(aq)
Ag(s)
AgNO3(aq)
GALVANIC CELL

## Slide 18: Electrochemistry Galvanic Cells
Description of how example cell works
  Reaction on anode = oxidation
  Anode = Zn electrode (as the Eº for Zn2+ is less than for that for Ag+)
  So, reaction on cathode must be reduction and involve Ag
  Oxidation produces e-, so anode has (–) charge (galvanic cells only); current runs from cathode to anode
  Salt bridge allows replenishment of ions as cations migrate to cathode and anions toward anodes
Salt Bridge
voltmeter
Zn(s)
ZnSO4(aq)
Ag(s)
AgNO3(aq)
GALVANIC CELL
Zn(s) → Zn2+ + 2e-
Ag+ + e- → Ag(s)
–
+

## Slide 19: Electrochemistry Galvanic Cells
Cell notation
  Example Cell:
  Zn(s)|ZnSO4(aq)||AgNO3(aq)|Ag(s)
Salt Bridge
voltmeter
Zn(s)
ZnSO4(aq)
Ag(s)
AgNO3(aq)
GALVANIC CELL
left side for anode (right side for cathode)
“|” means phase boundary
“||” means salt bridge