## Slide 1: Chem. 133 – 3/16 Lecture

## Slide 2: Announcements
Second Homework Set
Today’s Lecture
  Spectroscopy (Chapter 17)
    Region of Minimum Uncertainty (skipped last time)
  Spectroscopic Instrumentation (Chapter 19)

## Slide 3: Beer’s Law – Best Region for Absorption Measurements
Determine the best region for most precise quantitative absorption measurements if uncertainty in transmittance is constant
A
% uncertainty
0
2
High A values - Poor precision due to little light reaching detector
Low A values – poor precision due to small change in light

## Slide 4: Chapter 19 - Spectrometers
Main Components:
1) Light Source (produces light in right wavelength range)
2) Wavelength Descriminator (allows determination of signal at each wavelength)
3) Sample (in sample container)
4) Light Transducer (converts light intensity to electrical signal)
5 )Electronics (Data processing, storage and display)
Example: Simple Absorption Spectrophotometer
Light Source
(e.g tungsten lamp)
Monochromator
Sample
detector (e.g. photodiode)
Electronics
single l out

## Slide 5: Spectrometers
Some times you have to think creatively to get all the components.
Example NMR spectrometer:
Light source = antenna (for exciting sample, and sample re-emission)
Light transducer = antenna
Electronics = A/D board (plus many other components)
Wavelength descriminator =
Fourier Transformation
Radio Frequency Signal Generator
Antenna
A/D Board
Fourier Transformed Data

## Slide 6: Spectrometers – Fluorescence/Phosphorescence
Fluorescence Spectrometers
  Need two wavelength descriminators
  Emission light usually at 90 deg. from excitation light
  Can pulse light to discriminate against various emissions (based on different decay times for different processes)
  Normally more intense light and more sensitive detector than absorption measurements since these improve sensitivity
lamp
Excitation
monochromator
sample
Emission monochromator
Light detector

## Slide 7: Absorption Spectrometers
Sensitivity based on differentiation of light levels (P vs P0) so stable (or compensated) sources and detectors are more important
Dual beam instruments account for drifts in light intensity or detector response
Light Source
(tungsten lamp)
Monochromator
Sample
Electronics
chopper or beam splitter
Reference
detector

## Slide 8: Some Questions
Does the intensity of a light source have a large effect on the sensitivity of a UV absorption spectrometer? What about a fluorescence spectrometer?
If a sample is known to fluoresce and phosphoresce, how can you discriminate against one of these processes?
If a sample can both fluoresce and absorb light, why would one want to use a fluorescent spectrometer?
What is the advantage of using a dual beam UV absorption spectrometer?
List 5 components of spectrometers.
Why could the use of a broad band light source in the absence of wavelength discrimination lead to poor quantification of light absorbing constituents?

## Slide 9: Spectrometers – Specific Components Light Sources
Continuous Sources - General
  Provide light over a distribution of wavelengths
  Needed for multi-purpose instruments that read over range of wavelengths
  Sources are usually limited to wavelength ranges (e.g. D2 source for UV)

## Slide 10: Spectrometers – Light Sources
Continuous Sources – Specific
  For visible through infrared, sources are “blackbody” emitters
  For UV light, discharge lamps (e.g. deuterium) are more common (production of light through charged particle collision excitation)
  Similar light sources (based on charged particle collisions) are used for X-rays and for higher intensity lamps used for fluorescence
  For radio waves, light generated by putting AC signal on bare wire (antenna). Wide range of AC frequencies will produce a broad band of wavelengths.
UV
Vis
IR
high T
low T (max shifted to larger l)
intensity

## Slide 11: Spectrometers – Light Sources
Discrete Light Sources - General
  More common in “specific” instruments (e.g. industrial process instrument that measures single constituent)
  Light source usually is a (or the) wavelength discriminator also.
  Specific Sources
  LEDs (inexpensive light sources – relatively narrow band of wavelengths)
  Hollow cathode lamps (used in atomic absorption – discussed later)
  Lasers (intense, coherent, unidirectional, and very narrow wavelength distribution)

## Slide 12: Spectrometers – Wavelength Discrimination
Filters
  Mostly used with specific instruments
  “Standard Filters” – act to pass band of light or cut-off low or high wavelengths
  Interference filters
    pass a narrow band of light
    based on interference (show on board)
    used with other filters to reduce other orders
    some “tuning” of wavelength possible by changing gap or refractive index
intensity
before filter
after filter
intensity
before filter
wavelength
wavelength
after filter

## Slide 13: Spectrometers – Wavelength Discrimination
B. Monochromators
  Allows selection of a narrow band of wavelength from “broad band” source of light
  Most monochromators allow continuous adjustment of the selected wavelengths
  Some monochromators also allow adjustment of the range of wavelengths passed (Dl)
intensity
wavelength
after filter
before filter
desired l
Dl

## Slide 14: Spectrometers – Monochromators
A. Components
  Entrance Slit (to match exit slit)
  Light Collimator (optics to make light beam parallel when falling on dispersive element)
  Dispersing Element (to disperse light at different angles for different l values)
  Focusing Optics (to focus light on exit slit)
  Exit Slit (to select range of l values passed – Dl)
entrance slit
light
grating
collimating optics
l1
l2
Focusing optics
exit slit
In this example, wavelength selection occurs through rotation of the grating

## Slide 15: Spectrometers – Monochromators
B. Dispersion of Light
  Prisms – based on refractive index (n) = f(l)
  Gratings – based on constructive interference
    2 beams hitting grating will travel different distances
    travel difference = a – b
    this difference must be an integral # of l to lead to constructive interference
    a – b = nl (n = integer)
    from geometry, nl = d(sinq – sinf)
    Each groove acts as a light source
extra distance traveled by beam 2 = a
1
2
extra distance traveled by beam 1 = b
d
q
f
d = groove spacing
q = incoming light angle
f = outgoing light angle

## Slide 16: Spectrometers – Monochromators
B. Performance of Grating
  Resolution = l/Dl = nN
  where n = order (1, 2, 3...) and N = No. grooves illuminated
  To increase resolution,
    a. decrease d (groove spacing)
    increase length of grating illuminated (perpendicular to grooves)
    use higher diffraction order (n = 5 vs. n = 1)
  Dispersion from gratings:
    Angular dispersion = Df/Dl = n/dcosf
    Linear dispersion = D = Dy/Dl = FDf/Dl
f
Exit slit
y-axis
F = focal length

## Slide 17: Spectrometers – Monochromators
B. More on Linear Dispersion
  Dy = slit width = W: related to band width passed through monochromator (Dl)
  Dl = Wdcosf/Fn
  For better resolutions,
    Decrease W
    Use smaller d
    Use larger f
    Use larger F
    Use larger n
  All have drawbacks:
    a), c) and e) decrease light throughput
    b) Gratings more readily damaged
    d) Means larger monochromator
    e) Has more interferences from other n values

## Slide 18: Wavelength Discrimination Monochromators
Other Performance Measures (besides resolution)
  light throughput (% of light entering monochromator which exits monochromator)
  scanning range (λmin to λmax)
  stray light (light passed through monochromator outside of Δλ)

## Slide 19: Spectrometers Some Questions I
List one type of discrete light source.
List one method to create monochromatic light from a white light source without a monochromator.
List the five major components of a monchromator.