## Slide 1: Chem. 133 – 3/28 Lecture

## Slide 2: Announcements
Second Homework Set – Additional Problems due Thursday
Next Quiz on Thursday
Today’s Lecture
  Spectroscopic Instrumentation (Chapter 19)
    Wavelength Discrimination
      Filters (covered last time)
      Monochromators
      Polychromators
      Other methods
    Light Detectors
      Transducers
      Energy dispersive detectors

## Slide 3: Spectrometers – Monochromators
A. Components
  Entrance Slit (to match exit slit)
  Light Collimator (optics to make light beam parallel when falling on dispersive element)
  Dispersing Element (to disperse light at different angles for different l values)
  Focusing Optics (to focus light on exit slit)
  Exit Slit (to select range of l values passed – Dl)
entrance slit
light
grating
collimating optics
l1
l2
Focusing optics
exit slit
In this example, wavelength selection occurs through rotation of the grating

## Slide 4: Spectrometers – Monochromators
B. Dispersion of Light
  Prisms – based on refractive index (n) = f(l)
  Gratings – based on constructive interference
    2 beams hitting grating will travel different distances
    travel difference = a – b
    this difference must be an integral # of l to lead to constructive interference
    a – b = nl (n = integer)
    from geometry, nl = d(sinq – sinf)
    Each groove acts as a light source
extra distance traveled by beam 2 = a
1
2
extra distance traveled by beam 1 = b
d
q
f
d = groove spacing
q = incoming light angle
f = outgoing light angle

## Slide 5: Spectrometers – Monochromators
B. Performance of Grating
  Resolution = l/Dl = nN
  where n = order (1, 2, 3...) and N = No. grooves illuminated
  To increase resolution,
    a. decrease d (groove spacing)
    increase length of grating illuminated (perpendicular to grooves)
    use higher diffraction order (n = 5 vs. n = 1)
  Dispersion from gratings:
    Angular dispersion = Df/Dl = n/dcosf
    Linear dispersion = D = Dy/Dl = FDf/Dl
f
Exit slit
y-axis
F = focal length

## Slide 6: Spectrometers – Monochromators
B. More on Linear Dispersion
  Dy = slit width = W: related to band width passed through monochromator (Dl)
  Dl = Wdcosf/Fn
  For better resolutions,
    Decrease W
    Use smaller d
    Use larger f
    Use larger F
    Use larger n
  All have drawbacks:
    a), c) and e) decrease light throughput
    b) Gratings more readily damaged
    d) Means larger monochromator
    e) Has more interferences from other n values

## Slide 7: Wavelength Discrimination Monochromators
Other Performance Measures (besides resolution)
  light throughput (% of light entering monochromator which exits monochromator)
  scanning range (λmin to λmax)
  stray light (light passed through monochromator outside of selected Δλ)

## Slide 8: Spectrometers Some Questions I
List one type of discrete light source.
List one method to create monochromatic light from a white light source without a monochromator.
List the five major components of a monchromator.

## Slide 9: Spectrometers Some Questions II
If white light enters the monochromator to the right, which wavelength is longer wavelength?
List two parameters that will affect the resolution. Can any of these be easily changed?
A band pass filter is often placed between the grating and the focusing optics. What is the purpose of this filter?
If a grating is used with 320 lines/mm and the output angle for 380 nm is 45° and the focal length is 40 cm for 1st order light, what exit slit width is needed to be able to obtain a resolution of 200?
l1
l2
exit slit

## Slide 10: Spectrometers – Wavelength Discrimination
Polychromators
  1. In place of exit slit, an array of detectors exists
  This allows simultaneous recording of absorption over wavelength range
  No rotation of grating is needed
  Resolution (mainly) determined by width of detector element
  Dy = kDl
light
l1
l2
sample
Detector array top view
Detector element
Dy

## Slide 11: Spectrometers – Wavelength Discrimination
2-D Polychromators
  1. Light can be dispersed in two dimensions by placing a prism in front of the grating (dispersion in and out of the screen) to go along with the grating’s dispersion (in y-axis)
  See Color Plate 25 in Harris
  Requires 2-D detector array
  Usually uses high order grating dispersion (e.g. n = 11, 12, 13, 14) with different orders separated by prism
l1
l2
prism
2-D detector array
prism
dispersion
grating dispersion (y-axis)
emission light source
Detector elements

## Slide 12: Spectrometers – Wavelength Discrimination
D. Other Methods
  Energy-dispersive detectors (X-ray and g-ray analysis) – wavelength discrimination is part of detection system
  Fourier-transform Instruments
    Will cover for IR (today) and NMR
    “White” light passed through sample
    Variance in response with time or with distance is recorded and then transformed to conventional spectrum

## Slide 13: Wavelength Discrimination Fourier Transform Instruments
FTIR Instruments
  Uses Michelson interferometer (see Figure)
  Light goes to beam splitter (partially reflecting/partially transmitting
  Part of beam goes to fixed mirror and is reflected. Part of this beam then goes through the sample to the detector
  Another part of the original beam goes through the beam splitter to a moving mirror and is reflected with part of this going on to the sample and detector
light
Beam splitter
Fixed mirror
Mirror on drive
sample
detector

## Slide 14: Wavelength Discrimination Fourier Transform Instruments
FTIR Instruments (continued)
  If beams from the two paths combine “in phase” (both wave maxima) constructive interference occurs and greater light intensity reaches sample/detector
  If beams are not “in phase”, less light reaches detector
  Distance between beam splitter and mirror affects whether light is in phase
  Since “white” light is used (actually broad band IR), at different distances, different wavelengths will be in phase
  Recorded signal is Fourier transformed so plot of intensity vs. mirror distance or time is converted to intensity vs. frequency
intensity
Mirror position (or time if mirror moves)
l1
l2

## Slide 15: Wavelength Discrimination Fourier Transform Instruments
Performance:
  Δṽ (range of wavenumbers passed) is inversely related to distance traveled by mirror (D) (not explained clearly in text)
  This means better resolution (larger ṽ/Δṽ) when D is larger
  Spectral range depends on sampled data speed (assuming fast detector)
  High resolution over a long wavenumber range will take more time
small displacement → poor resolution

## Slide 16: Spectrometers – Light Detectors
Detectors covered in electronics section
  UV/Vis/NearIR: Photocell, photomultiplier tube, photodiode, photoconductivity cell, and solid state array detectors (charged coupled device or CCD)
  IR: temperature measurement (e.g. thermopile), and solid state
  NMR: antenna

## Slide 17: Spectrometers – Light Detectors
Detectors for high energy (X-ray, g-ray light) (both gas cells and solid state available)
  Due to high energy, a single photon can easily produce a big signal
  Two types: gas cells (e.g. Geiger Counter) and solid state sensors (e.g. Si(Li) detectors)
  In both cases, detectors can be set up where cascade of electrons is produced from a single photon
  The number of ions produced from photons can be dependent upon the photon energy
time
current
high E photon
low E photon
energy
counts/s
solid state detector
I
+
+
+
-
-
-
These detectors are said to be energy dispersive (no monochromator needed)