## Slide 1: Chem. 133 – 3/30 Lecture

## Slide 2: Announcements
Second Homework Set – due today (additional problems); key will be posted soon
Quiz Today
Today’s Lecture
  Spectroscopic Instrumentation (Chapter 19)
    Light Detectors
      Transducers
      Energy dispersive detectors
  Atomic Spectroscopy (Chapter 20)
    Overview + methods for solids (not in text)
    Theory
    Atomization with flames (if time)

## Slide 3: Spectrometers – Light Detectors
Detectors covered in electronics section
  UV/Vis/NearIR: Photocell, photomultiplier tube, photodiode, photoconductivity cell, and solid state array detectors (charged coupled device or CCD)
  IR: temperature measurement (e.g. thermopile), and solid state
  NMR: antenna

## Slide 4: Spectrometers – Light Detectors
Detectors for high energy (X-ray, g-ray light) (both gas cells and solid state available)
  Due to high energy, a single photon can easily produce a big signal
  Two types: gas cells (e.g. Geiger Counter) and solid state sensors (e.g. Si(Li) detectors)
  In both cases, detectors can be set up where cascade of electrons is produced from a single photon
  The number of ions produced from photons can be dependent upon the photon energy
time
current
high E photon
low E photon
energy
counts/s
solid state detector
I
+
+
+
-
-
-
These detectors are said to be energy dispersive (no monochromator needed)

## Slide 5: Atomic Spectroscopy Overview
Main Purpose
  Determine elemental composition (or concentration of specific elements)
Main Performance Concerns
  Sensitivity
  Multi-element vs. single element
  List of useful elements (most methods work well with most metals, poorly with non-metals)
  Speed (instrument plus sample preparation)
  Interferences (for different matrices)
  Precision
  Required sample preparation

## Slide 6: Atomic Spectroscopy Overview
Instrument Types
  Analysis for liquid samples (main focus of text + lecture discussion)
  Systems for solid samples
    Modified instruments for liquids (involving conversion to gas phase first)
      2 examples in book: graphite furnace with solid sample placed in tube (see p. 485) and laser ablation (see p. 495)
      laser ablation allows microanalysis
    X-ray Fluorescence Spectroscopy and X-ray Emission Detection attachments coupled to electron microscopy
      Both based on spectral (or energy-dispersive) analysis of emitted X-rays to determine elements present

## Slide 7: Atomic Spectroscopy Overview
Instrument Types – Systems for Solids – cont.
  XRF – cont.
      Emitted X-rays have wavelengths dependent upon element (but generally not element’s charge or surroundings)
      Accurate quantification is more difficult due to limited penetration of sample by X-rays or electrons and by attenuation of emitted X-rays due to absorption (matrix effects)
      Sensitivity and selectivity somewhat less than standard methods

## Slide 8: Atomic Spectroscopy Overview
Instrument Types – For Analysis of Liquids
  Atomization Systems: to convert elements to gaseous atoms or ions (MS detection)
    Flame
    Electrothermal (Graphite Furnace)
    Inductively Coupled Plasma (ICP)
  Atom Detection: to detect atoms (or ions in MS)
    Atomic Absorption Spectroscopy (with flame or electrothermal)
    Atomic Emission Spectroscopy (mainly with ICP)
    Mass Spectrometry (with ICP) – only detects ions

## Slide 9: Atomic Spectroscopy Theory
Spectroscopy is performed on atoms in gas phase
Transitions are very simple (well defined energy states with no vibration/rotation /solvent interactions)
Allowed transitions depend on selection rules (not covered here)
E
Na(g)o (3s)
4s
4p
5s
5p
absorption

## Slide 10: Atomic Spectroscopy Theory
Consequence of well defined energy levels:
  very narrow absorption peaks
  few interferences from other atoms
  very good sensitivity (all absorption occurs at narrow l range)
  but can not use standard monochromator where Dl (from monochromator) >> dl due to extreme deviations to Beer’s law
  requires greater wavelength discrimination for absorption measurements
l
A
Spectrum from high resolutions spectrometer (not typical for AA)
atomic transition
molecular transition
very narrow natural peak width (dl ~ 0.001 nm)
broader width

## Slide 11: Atomic Spectroscopy Theory
For emission measurements, a key is to populate higher energy levels
In most cases, this occurs through the thermal methods also responsible for atomization
Fraction of excited energy levels populated is given by Boltzmann Distribution
More emission at higher temperatures and for longer wavelengths (smaller DE)
Na(g)o (3s)
4p
E
N = number atoms in ground (0) and excited (*) states
g = degeneracy (# equivalent states) = 3 in above example (for g*); 1 for g0
k = Boltzmann constant = 1.38 x 10-23 J/K

## Slide 12: Atomic Spectroscopy Theory
Example problem:
Calcium absorbs light at 422 nm. Calculate the ratio of Ca atoms in the excited state to the ground state at 3200 K (temperature in N2O fueled flame). g*/g0 = 3 (3 5p orbitals to 1 4s orbital).
Notes: 12

## Slide 13: Atomic Spectroscopy Atomization
Flame Atomization
  used for liquid samples
  liquid pulled by action of nebulizer
  nebulizer produces spray of sample liquid
  droplets evaporate in spray chamber leaving particles
  fuel added and ignited in flame
  atomization of remaining particles and spray droplets occurs in flame
  optical beam through region of best atomization
sample in
fuel (HCCH)
oxidant (air or N2O)
burner head
spray chamber
nebulizer
light beam
nebulizer
air
liquid

## Slide 14: Atomic Spectroscopy Atomization
Atomization in flames – Processes
  nebulization of liquid: MgCl2(aq) → MgCl2(spray droplet)
  evaporation of solvent: MgCl2(spray droplet) → MgCl2(s)
  Volatilization in flame: MgCl2(s) → MgCl2(g)
  Atomization (in hotter part of flame): MgCl2(g) → Mg(g) + Cl2(g)
Target species for absorption measurement

## Slide 15: Atomic Spectroscopy Atomization
Complications/Losses
  Ideally, every atom entering nebulizer ends up as gaseous atom
  In practice, at best only a few % of atoms become atoms in flame
  The nebulization process is not that efficient (much of water hits walls and goes out drain)
  Poor volatilization also occurs with less volatile salts (e.g. many phosphates)

## Slide 16: Atomic Spectroscopy Atomization
Complications/Losses (continued)
  Poor atomization also can occur due to secondary processes such as:
    Formation of oxides + hydroxides (e.g. 2Mg (g) + O2 (g) → 2MgO (g))
    Ionization (Na (g) + Cl (g) → Na+ (g) + Cl- (g))
  If the atomization is affected by other compounds in sample matrix (e.g. the presence of phosphates), this is called a matrix effect (discussed more later)