## Slide 1: Chem. 133 – 4/4 Lecture

## Slide 2: Announcements
Second Homework Set – (additional problem – slight error in first 2.1.1 key posted; since fixed)
Today’s Lecture
  Atomic Spectroscopy (Chapter 20)
    Theory (Boltzmann Distribution Problem)
    Atomization
      Flame
      Electrothermal
      ICP
    Spectrometers
      AA instruments
      AE instruments

## Slide 3: Atomic Spectroscopy Theory
Example problem:
Calcium absorbs light at 422 nm. Calculate the ratio of Ca atoms in the excited state to the ground state at 3200 K (temperature in N2O fueled flame). g*/g0 = 3 (3 5p orbitals to 1 4s orbital).
Notes: 3

## Slide 4: Atomic Spectroscopy Atomization
Flame Atomization
  used for liquid samples
  liquid pulled by action of nebulizer
  nebulizer produces spray of sample liquid
  droplets evaporate in spray chamber leaving particles
  fuel added and ignited in flame
  atomization of remaining particles and spray droplets occurs in flame
  optical beam through region of best atomization
sample in
fuel (HCCH)
oxidant (air or N2O)
burner head
spray chamber
nebulizer
light beam
nebulizer
air
liquid

## Slide 5: Atomic Spectroscopy Atomization
Atomization in flames – Processes
  nebulization of liquid: MgCl2(aq) → MgCl2(spray droplet)
  evaporation of solvent: MgCl2(spray droplet) → MgCl2(s)
  Volatilization in flame: MgCl2(s) → MgCl2(g)
  Atomization (in hotter part of flame): MgCl2(g) → Mg(g) + Cl2(g)
Target species for absorption measurement

## Slide 6: Atomic Spectroscopy Atomization
Complications/Losses
  Ideally, every atom entering nebulizer ends up as gaseous atom
  In practice, at best only a few % of atoms become atoms in flame
  The nebulization process is not that efficient (much of water hits walls and goes out drain)
  Poor volatilization also occurs with less volatile salts (e.g. many phosphates)

## Slide 7: Atomic Spectroscopy Atomization
Complications/Losses (continued)
  Poor atomization also can occur due to secondary processes such as:
    Formation of oxides + hydroxides (e.g. 2Mg (g) + O2 (g) → 2MgO (g))
    Ionization (Na (g) + Cl (g) → Na+ (g) + Cl- (g))
  If the atomization is affected by other compounds in sample matrix (e.g. the presence of phosphates), this is called a matrix effect (discussed more later)

## Slide 8: Atomic Spectroscopy Atomization
Electrothermal Atomization
  Atomization occurs in a graphite furnace
  Process is different in that a small sample is placed in a graphite tube and atomization occurs rapidly but in a discontinuous manner
  Electrothermal atomization is more efficient; atoms spend more time in the beam path, and less sample is required resulting in much greater sensitivity
    Concentration LODs are typically ~100 times lower (e.g. 100 ppt for EA vs. 10 ppb for flame)
    Mass LODs are even lower (100 pg/mL*0.01 mL = 1 pg for EA vs. 10 ng/mL*2 mL = 20 ng for flame)

## Slide 9: Atomic Spectroscopy Atomization
Electrothermal Atomization (Process)
  Sample is placed through hole onto L’vov platform
  Graphite tube is heated by resistive heating
  This occurs in steps (dry, char, atomize, clean)
Graphite Tube in Chamber (not shown)
L’vov Platform
Sample in
T
time
dry
char
atomize
Clean + cool down
Ar in chamber flow stops and optical measurements made

## Slide 10: Atomic Spectroscopy Atomization
Inductively Coupled Plasma (ICP)
  A plasma is induced by radio frequency currents in surrounding coil
  Once a spark occurs in Ar gas, some electrons leave Ar producing Ar+ + e-
  The sample is introduced by nebulization in the Ar stream
  The accelerations of Ar+ and e- induce further production of ions and great heat production
  Much higher temperatures are created (6000 K to 10000 K vs. flames)
ICP Torch
Quartz tube
Argon + Sample
RF Coil
Plasma

## Slide 11: Atomic Spectroscopy Atomization
Advantages of ICP Atomization
  Greater atomization efficiency than in flame AA (partly because better nebulizers are used than with flames due to higher total instrument cost and partly due to higher temperatures)
  Fewer matrix effects because atomization is more complete at higher temperatures
  High temperature atomization allows much greater emission flux + more ionization allowing coupling with emission spectrophotometers and mass spectrometers
  Emission and MS allow faster multi-element analysis

## Slide 12: Chapter 20 Questions
Why would it be difficult to use a broadband light source and monochromator to produce light used in AA spectrometers?
List three methods for atomizing elements.
List two processes that can decrease atomization efficiency in flame atomization.
What is an advantage in using electrothermal atomization in AAS?
Which atomization method tends to result in the most complete breakdown of elements to atoms in the gas phase?
Why is ICP better for emission measurements than flame?

## Slide 13: Atomic Spectroscopy Absorption Spectrometers
The lamp is a hollow cathode lamp containing the element(s) of interest in cathode
The lamp is operated under relatively cool conditions at lower pressures to reduce Doppler and pressure broadening of atomic emission lines
A very narrow band of light emitted from hollow cathode lamps is needed so that absorption by atoms in flame mostly follows Beer’s law
The monochromator serves as a coarse filter to remove other wavelength bands from light and light emitted from flames
Lamp source
Flame or graphite tube
monochromator
Light detector

## Slide 14: Atomic Spectroscopy Absorption Spectrometers
A narrower emission spectrum from hollow cathode lamp (vs. flame absorption) results in better Beer’s law behavior
wavelength
Intensity or absorbance
hollow cathode lamp emission
Atomic absorption spectrum in flame
Additional broadening in flame from temperature (Doppler) or pressure

## Slide 15: Atomic Spectroscopy Interference in Absorption Measurements
Spectral Interference
  Very few atom – atom interferences
  Interference from flame (or graphite tube) emissions are reduced by modulating lamp
    no lamp: signal from flame vs. with lamp
    then with lamp: signal from lamp + flame – absorption by atoms
  Interference from molecular species absorbing lamp photons (mostly at shorter wavelengths and light scattering in EA-AA)
  This interference can be removed by periodically using a deuterium lamp (broad band light source) or using the Zeeman effect (magnetic splitting of absorption bands)

## Slide 16: Atomic Spectroscopy Interference in Absorption Measurements
Chemical Interference
  Arises from compounds in sample matrix or atomization conditions that affects element atomization
  Some examples of specific problems (mentioned previously) and solutions:
    Poor volatility due to PO43- – add Ca because it binds strongly to PO43- allowing analyte metal to volatilize better or use hotter flames
    Formation of metal oxides and hydroxides – use fuel rich flame
    Ionization of analyte atoms – add more readily ionizable metal (e.g Cs)
  Another approach is to use a standard addition calibration procedure (this won’t improve atomization but it accounts for it so that results are reliable)

## Slide 17: Atomic Spectroscopy Interference in Absorption Measurements
Standard Addition
  Used when sample matrix affects response to analytes
  Commonly needed for AAS with complicated samples
  Standard is added to sample (usually in multiple increments)
  Needed if slope is affected by matrix
  Concentration is determined by extrapolation (= |X-intercept|)
Absorbance
Concentration Added
Analyte Concentration
standards in water
Sample

## Slide 18: Atomic Spectroscopy Emission Spectrometers
In emission measurements, the plasma (or flame) is the light source
Flame sources are generally limited to a few elements (only hot enough for low E – visible light emissions)
A monochromator or polychromator is the means of wavelength discrimination
Sensitive detectors are needed
ICP-AES is faster than AAS because switching monochromator settings can be done faster than switching lamp plus flame conditions
Plasma (light source + sample)
Monochromator or Polychromator
Light detector or detector array
Liquid sample, nebulizer, Ar source

## Slide 19: Atomic Spectroscopy Emission Spectrometers
Sequential vs. Simultaneous Instruments
Sequential Instruments use:
  A standard monochromator
  Select for elements by rotating the monochromator grating to specific wavelengths
Simultaneous Instruments use:
  A 1D or 2D polychromator (Harris Color Plate 24/25)
  1D instruments typically use photomultiplier detectors behind multiple exit slits
  2D instrument shown in 4/1 lecture slide 13
  Selected elements (1D instruments) or all elements can be analyzed simultaneously resulting in faster analysis and less sample consumption.

## Slide 20: Atomic Spectroscopy Interference in Emission Measurements
Interferences
  Atom – atom interferences more common than in atomic absorption because monochromators offer less selectivity than hollow cathode lamps
  Interference from molecular emissions are reduced by scanning to the sides of the atomic peaks
  Chemical interferences are less prevalent due to greater atomization efficiency
Emission Spectrum
Atomic peak
background