## Slide 1: Chem. 133 – 4/6 Lecture

## Slide 2: Announcements
Today’s Lecture
  Atomic Spectroscopy (Chapter 20)
    Spectrometers
      AA instruments (last time)
      AE instruments
      Mass Spectrometers (briefly)
  NMR Spectroscopy (Rubinson & Rubinson Ch. 11)
    Theory

## Slide 3: Atomic Spectroscopy Emission Spectrometers
In emission measurements, the plasma (or flame) is the light source
Flame sources are generally limited to a few elements (only hot enough for low E – visible light emissions)
A monochromator or polychromator is the means of wavelength discrimination
Sensitive detectors are needed
ICP-AES is faster than AAS because switching monochromator settings can be done faster than switching lamp plus flame conditions
Plasma (light source + sample)
Monochromator or Polychromator
Light detector or detector array
Liquid sample, nebulizer, Ar source

## Slide 4: Atomic Spectroscopy Emission Spectrometers
Sequential vs. Simultaneous Instruments
Sequential Instruments use:
  A standard monochromator
  Select for elements by rotating the monochromator grating to specific wavelengths
Simultaneous Instruments use:
  A 1D or 2D polychromator (Harris Color Plate 24/25)
  1D instruments typically use photomultiplier detectors behind multiple exit slits
  2D instrument shown in 4/1 lecture slide 13
  Selected elements (1D instruments) or all elements can be analyzed simultaneously resulting in faster analysis and less sample consumption.

## Slide 5: Atomic Spectroscopy Interference in Emission Measurements
Interferences
  Atom – atom interferences more common than in atomic absorption because monochromators offer less selectivity than hollow cathode lamps
  Interference from molecular emissions are reduced by scanning to the sides of the atomic peaks
  Chemical interferences are less prevalent due to greater atomization efficiency
Emission Spectrum
Atomic peak
background

## Slide 6: Atomic Mass Spectrometry
Most common arrangement consists of ICP torch leading to MS interface
The Ar+ ions (and electrons) collide with metals leading to ionization
The MS interface consists of skimmer cones to allow ions in, and to drop the pressure in stages, and ion optics (covered when we cover MS)
ICP-MS typically is the most sensitive elemental analysis method
Interference can arise from metals (e.g. 138Ba2+ vs. 69Ga+) or from ICP species (e.g. 40Ar+ and 40Ca+)
Use of secondary isotopic masses and collision cell reactions can reduce these interferences
Plasma (atomizer + ion source)
Liquid sample, nebulizer, Ar source
Mass spectrometer (e.g. quadrupole)
collision cell

## Slide 7: Atomic Spectroscopy Comparison of Instruments
Instrument | Cost | Speed | Sensitivity
Flame-AA | Low (~$10-15K) | Slow | Moderate (~0.01 ppm)
GF-AA | Moderate (~$40K) | Slowest | Very Good
Sequential ICP-AES | Moderate | Medium | Moderate
Simultaneous ICP-AES | High | Fast | Good
ICP-MS | Highest (~$200K) | Fast | Excellent

## Slide 8: Atomic Spectroscopy Some Questions
Why is AES with a plasma normally more sensitive than AES with a flame?
List two ways in which a process in a flame can lead to reduced sensitivity and a way to deal with each process so its effect on the analysis is minimized.
Why can a simultaneous ICP-AES be more sensitive than an sequential ICP-AES if used for analysis of 12 metals?
If a sample matrix produces molecular emissions that interfere with atomic emissions, how would this be observed and how can this be accounted for?
What can cause interferences in ICP-MS?

## Slide 9: Nuclear Magnetic Resonance (NMR) Spectrometry Major Uses
Identification of Pure Compounds (Qualitative Analysis)
Structural Determination (e.g. protein shape)
Quantitative Analysis
Characterization of Compounds in Mixtures (e.g. % of C as aromatic C)
Imaging (MRI) – not covered

## Slide 10: NMR Spectrometry Theory
Spin
  a magnetic property that sub atomic particles have (electrons, some nuclei)
  some combinations do not result in observable spin (paired electrons have no observable spin; many nuclei have no observable spin)
  Electron spin transitions occur at higher energies and are the basis of electron paramagnetic spectroscopy (EPR)
  Nuclear spin given by Nuclear Spin Quantum Number (I)

## Slide 11: NMR Spectrometry Theory
Nuclear Spin (continued)
  I = 0 nuclei → no spin (not useful in NMR) – e.g. 12C
  I = ½ nuclei → most commonly used nuclei (1H, 13C, 19F, many others)
  I > 1 nuclei → used occasionally, important for spin-spin coupling
  number of different spin states (m) = 2I + 1
  examples:
    1H (I = ½), 2 states
    2H (I = 1), 3 states
up state (m = +1/2)
down state (m = -1/2)
up state (m = 1)
middle state (m = 0)
down state (m = -1)

## Slide 12: NMR Spectrometry Theory
Effect of External Magnetic Field on Nuclei States
  aligned nuclei (m = +1/2) have slightly lower energy (are more stable) than anti-aligned states (m = -1/2)
  the greater the magnetic field (H), the greater the energy difference between the states
Applied Magnetic Field H0* (also have used B0 in past)
“up” state – m = +1/2
“down” state – m = -1/2
Note: arrows drawn at angles because spin vectors precess about H
path made by vector tips
*Note: technically H is the magnetic field at the nucleus which is not quite the same as the applied magnetic field (H0)

## Slide 13: NMR Spectrometry Theory
Energy depends on nucleus, spin state (m), and magnetic field
g (gamma) = magnetogyric ratio (constant for given nuclei) and h = Planck’s constant
Energy difference
Energy
H0
ΔE
m = +1/2
m = -1/2

## Slide 14: NMR Spectrometry Theory
Transitions between the ground and excited state can occur through absorption of light
Lowest Resolution Spectroscopy
CH3CF2OH
or
signal
Ho
1H
19F
13C (small because most C is 12C)
Ho is traditionally used for x-axis because older instruments involved changing Ho (most newer instrument don’t). A frequency plot at constant field would be reversed (1H at highest frequency).
H scanned at fixed n

## Slide 15: NMR Spectrometry Theory
Frequency depends on g and Ho.
Intensity (y-axis) depends on:
  ΔE (will cover later)
  number of nuclei in compound (e.g. for 13C1H4, there are 4 times as many Hs as Cs)
  isotopic abundance (e.g. for non-isotopically enriched organics, 13C is only ~1% of all C).
  other factors (e.g. relaxation times)

## Slide 16: NMR Spectrometry Theory
Effects of ΔE
  Opposite problem in Boltzmann distribution as in AES: too many nuclei in excited state
  ΔE is much smaller for NMR; e.g. ΔE for n = 300 MHz = 2.0 x 10-25 J (Ho = 70.5 kgauss) vs. for λ = 400 nm, ΔE = 5.0 x 10-19 J
  N*/N0 (n = 300 MHz, T = 298 K) = e-ΔE/kT = 0.999951
  Why is this a problem (especially when for AES too few excited states was a problem)?
  Problem occurs because absorption of light can only be observed if there is a difference in population of states.
Example: element with 3 nuclei in ground and 2 in excited state
Absorption of light: promotes nuclei
Induced emission: knocks excited nuclei out of excited state releasing extra photon
Net: 2 photons in + 2 photons out

## Slide 17: NMR Spectrometry Theory
Effects of ΔE, continued
  Can only “see” excess nuclei (as absorption and emission near balanced)
  Back to the case of 1H in a 300 MHz NMR
  N*/N0 = 0.999951; if 400,000 nuclei, 10 more in ground state (200,005 ground, 199,995 excited)
Sample before absorption
m = +1/2
m = -1/2
...
Absorption of light
up to 5 nuclei can flip spins
detectable excess
Now, sample is saturated (invisible), until nuclei return to ground state

## Slide 18: NMR Spectrometry Theory
Consequence of Limited Nuclei Available for Absorption
  Lack of sensitivity (only 5 out of 400,000 nuclei available for observation – combined with insensitivity of detecting radio waves)
  Repeating absorption experiments requires time for excited nuclei to return or “relax” to ground states
Decay Process
  Once saturation occurs, no further absorption can occur until excited nuclei return to ground state
  2 types of decay (or relaxation) processes occur:
    spin-lattice relaxation (through nuclei interaction with surrounding molecules)
    spin-spin relaxation (relaxation by flipping neighboring nuclei – but this doesn’t affect saturation problem)

## Slide 19: NMR Spectrometry Theory
Decay Process (continued)
  Relaxation affects:
    rate at making absorption measurements (fast decay is better)
    peak widths (through Heisenberg Uncertainty Principle)
    δEδt = h or δnδt = 1 or δn = 1/δt (δn = peak width and δt = decay time)
    So, fast decay results in broader peaks
    An example is solids where spin-spin relaxation is fast; broad peaks result despite not fast spin-lattice relaxation

## Slide 20: NMR Spectrometry Some Questions
Modern NMRs continuously monitor 2H absorbance to account for magnetic field drift in the “lock” unit. The frequency of the 2H signal is observed to drift by 30Hz over 1 hour. Given the magnetic field H0= 8.45 T, γ(2H) = 8.22 x 107 radian T-1 s-1 and γ(1H) = 2.68 x 108 radian T-1 s-1, calculate the magnetic field drift and the drift in the 1H frequency in an hour.
17O has an I value of 5/2. How many spin states will it have?
Explain why sensitivity is increased by going to a larger magnetic field.
Will increasing the temperature increase or decrease NMR sensitivity (assuming it has no effect on relaxation processes)?
Notes: 20